EML 4905 Senior Design Project

A b.s. tHESIS

Prepared in Partial Fulfillment of the

Requirement for the Degree of

Bachelor of Science

in

Mechanical Engineering

IMPLANTABLE MEDICAL MICROCONNECTOR

100% Final Report

James Ciurdar

Eric Doan

Samuel Grada

MariaMartha Lopez

Advisor: Professor W. Kinzy Jones

April 3, 2014

This B.S. thesis is written in partial fulfillment of the requirements in EML 4905. The contents represent the opinion of the authors and not the Department of Mechanical and Materials Engineering.

Ethics Statement and Signatures

The work submitted in this B.S. thesis is solely prepared by a team consisting of Eric Doan, James Ciurdar, Samuel Grada, and MariaMartha Lopez and it is original. Excerpts from others’ work have been clearly identified, their work acknowledged within the text and listed in the list of references. All of the engineering drawings, computer programs, formulations, design work, prototype development and testing reported in this document are also original and prepared by the same team of students.

Eric Doan

Team Leader

James Ciurdar

Team Member

Samuel Grada

Team Member

MariaMartha Lopez

Team Member

|  | Dr. W. Kinzy Jones Faculty Advisor |  |
| --- | --- | --- |

Table of Contents

Chapter	Page

Cover Page	  i

Ethics Statement and Signatures	ii

Table of Contents	iii

List of Figures	v

List of Tables	vi

Abstract	1

1. Introduction	2

1.1 Problem Statement	2

1.2 Motivation	2

1.3 Literature Survey	3

1.3.1 Current products.	3

1.3.2 Biocompatible materials.	4

1.3.3 Lead attachment.	6

1.3.4 O-rings.	7

1.3.5 Devices used to manufacture.	8

2. Design Alternatives	11

2.1 Overview of Conceptual Design	11

2.2 Design Alternate 1	12

2.3 Design Alternate 2	13

2.4 Proposed Design	14

3. Project Management	15

3.1 Breakdown of Work into Specific Tasks	15

3.2 Organization of Work and Timeline	17

4. Engineering Design and Analysis	18

4.1 Structural Design	18

5. Prototype Construction	18

5.1 Description of Prototype	18

5.2 Prototype Cost Analysis	20

5.3 Prototype Iterations	21

6. Design Considerations	24

6.1 Design Adaptation due to Manufacturability	24

6.2 Interconnection Options	36

6.3 Environmental Impact	40

7. Testing and Evaluation	41

7.1 Overview	41

7.2 Final Design Cost Evaluation	43

8. Conclusion	44

8.1 Conclusion and Discussion	44

9. References	46

LIST OF FIGURES

Figure	Page

- Typical LTCC Manufacturing Process Schematic. (Vasudev, 2012)	5

- Common Problems Encountered During LTCC	5

- O-Ring Pressure. (McGrosky, 2013)	7

- Prototype Trial I	10

- Close-Up View of Microconnector.	10

- Preliminary Design of Microconnector.	12

- Circular Microconnector Design Alternative.	13

- Proposed Design - Pins and O-Ring Placement.	14

- Proposed Design - Closed and Sealed.	15

- Prototype Design in Comparison to a Dime.	19

- Cost for Half of a Prototype.	21

- Prototype Design with Carbon Tape as Sacrificial Material. 	23

- Close Up View of Silver Pin in LTCC Layers.	23

- Former Design using Screen Printing	24

- Male Half of New Design.	25

- Female Half of New Design.	25

- Former Design with the O-ring.	26

- Housing Created and Ceramic Part Placed in the Housing.	26

- Isometric View of Housing Shown with Alignment Pins	27

- Top View of Alignment Pins in Housing.	28

- View of the Housing with Alignment Pins Before Connection.	28

- Larger 3D Printed Half of Housing.	30

- Size Comparison of 3D Printed Housings.	30

- Measurement of Housing.	31

- Isometric View of the Microconnector.	31

- Top View of Microconnector.	32

- Front View of Microconnector.	32

- Addition of Strain Relief to the Housing.	33

- Fully Assembled Female Half of the Microconnector.	33

- Fully Assembled Male Side of Microconnector.	34

- Fully Assembled Front View of Device.	34

- Fully Assembled Isometric View of Part.	34

- Exploded View of How Microconnector will be Assembled.	35

- Major Dimensions of Microconnector.	35

- Fuzz Button Under Magnification.	37

- Fuzz Button Compared to Dime.	37

- Fuzz Button Comparison to Female Vias.	38

- Male and Female Half with Fuzz Button.	40

- Waste Material Compared to a Quarter.	41

- Projected Cost of Microconnector.	44

LIST OF TABLES

Table	Page

- Work Distribution.	16

- Project Timeline for Fall 2013.	17

- Updated Project Timeline from Spring 2014.	17

- Cost Analysis.	21

- Fixed Conditions for Fine Leak Test.	43

# Abstract

Implantable neurostimulation devices generate electrical stimulation to treat various disorders involving neurological problems. A neurostimulator is composed of a connector assembly which connects to the leads, thin wires that deliver electrical pulses from the neurostimulator. There currently are a number of neurostimulators treating a wide range of disorders including migraines, back pain, hearing loss, epilepsy, and Parkinson’s disease.  These devices are hard-wired to the leads and have a limited number of electrodes.   The permanent attachment of the leads to the neurostimulator means that failure from either component will result in removal of the entire device as opposed to a simple disconnection and replacement of one component with the use of a connector.  The leads must be hard-wired because there is no connector commercially available.

Neurostimulators currently have up to 16 electrodes. While they may be effective in treating certain disorders, neurostimulators are inadequate to work with neural prostheses being developed as well as the treatment of more advanced neurological disorders such as Alzheimer’s. The objective of this project is to develop a hermetic, 32 channel, implantable microconnector that is compatible for neurostimulator devices in the market. The connector is made with High Temperature Co-fired Ceramic (HTCC), a common material used in microelectronics along with titanium and platinum.  The dimensions of the initial prototype are 5mm by 4mm by 3mm. Overall, this electronic packaging will provide an easier way to assemble the neurostimulator in addition to improving neural feedback.

# 1. Introduction

# 1.1 Problem Statement

Implanted neurostimulation devices employ a hardwired design, which means that the stimulating pads and leads are hardwired directly to the generator device. While this design is functional, it leaves little room for improvement and for the exchange of devices. In the case where the generator device must be changed, which may be due to various reasons such as generator failure, the entire assembly must be detached and taken out of the body. The main problem with this approach is the human body’s tendency to adapt to foreign material. When the pads and leads are initially implanted, they are attached to the nerves or tissue, and the nerves begin to grow around the pads themselves. Therefore, tearing the tissue or nerve in order to detach the pads is very damaging and may harm the nerves or tissue.

Current device connections do not accommodate devices which require more stimulation leads than others. The treatment of neurological diseases such as Parkinson’s disease, Lou Gehrig’s disease, or Alzheimer’s disease require more stimulation leads simply because current research cannot pinpoint the exact locations of the affected areas in the brain. Because of this, more leads are used to solve the problem with minimal trial and error in the placement of the leads.

# 1.2 Motivation

The motivation for this project is based on providing a more feasible way to interchange neurostimulation devices. Creating an advanced connector improves the task for the surgeon in the operating room when connecting every lead to the correct nerve. Research done at St. Jude’s Children’s Hospital has led to new devices being implanted which help children with various physical and neurological problems. Because technology in the medical field advances so quickly, a need for a quick, easy, reliable connection is necessary to help push research to its fullest potential. The replacement of failed devices that were implanted into hundreds of children could have devastating effects on the recovery of children that never had the ability to see or hear.

Another application outside of research would be to use the connector to help rehabilitate retired veterans who were injured while on active duty. Many veterans come back home with no hope of ever walking again or having full use of one or more of their limbs. This connector has the potential to aid in Veterans Affairs’ research of new devices. Devices built for the stimulation of a prosthetic limb or treatment of non-ordinary neurological diseases may be more involved than traditional implants, thus an increased amount of leads connected to the nerves and tissues are necessary.

# 1.3 Literature Survey

# 1.3.1 Current products.  The survey began by investigating what is currently available in the marketplace. At this time, there is a patent issued to Neurostream Technologies for a high density implantable connector for neurostimulators.  This device has 9-channel connectors.  The designs were sold to a German company called Ottobock.  Ottobock is using this device exclusively for their stimulator to treat stroke victims suffering from foot drop.  Currently, this connector is not available as a standalone product.  The second problem with this product is that it only has nine channels.  Many neurostimulators in the market have 16 channels.  Moreover, the size of this product poses a problem.  While the patent application did not provide the dimensions, their scale appears much larger than our proposed design of 65 mils by 16 mils by 16 mils.  This gives the physician more flexibility in the placement of the connector.

# 1.3.2 Biocompatible materials. This device is implanted into the human body; consequently, the materials used for this device case must be biocompatible.  There are a number of materials that are compatible with the human body ranging from certain metals, some plastics, and some ceramics.

HTCC is High Temperature Co-fired Ceramic which is biocompatible while LTCC is Low Temperature Co-fired Ceramic which is not biocompatible.  These materials are widely used in microelectronics. Preliminary research indicates that LTCC and silver are used for the internal microelectronics. These materials play a vital role in building a device so small with many connectors.

Some metals that are biocompatible include stainless steels, cobalt alloys, and titanium alloys.  The focus was on titanium as this metal has a high tensile strength, low density and is highly resistant to bodily fluids.  Initially, titanium was the primary material used for neurostimulator and pacemaker cases.

The following is a schematic outlining a typical device manufacturing process using LTCC material, followed by examples of common problems encountered when working with LTCC. Common problems which are known to occur include “potato-chipping,” which is what may happen when the specimen shrinks at uneven rates during co-firing. Another common problem is the collapse of channels or vias. It is common practice to combat the collapse of desired channels or vias by filling the empty spaces with a sacrificial material before lamination. This sacrificial material, such as graphite tape or paste is then burned off in the sintering process. The following images were taken from Vasudev, 2012.

Figure 1: Typical LTCC Manufacturing Process Schematic. (Vasudev, 2012)

Figure 2: Common Problems Encountered During LTCC

Component Manufacturing. (Vasudev, 2012)

# 1.3.3 Lead attachment. Conventional neurostimulator systems consist of leads connected directly from the nerves through an extension (if necessary) and subsequently into the neurostimulator. There are a few methods of attaching leads depending on what type of leads will be used. Proximal leads, which are the lead ends connected directly to the neurostimulator, are often set into the connection junction using setscrews once the leads are set into position. As described in the proposed design of the connector, the connection points lead to corresponding pads to which the incoming lead ends will be connected. The employment of this connection system requires each half of the connector to link to the corresponding leads ends (one half to the stimulator, the other to the extension leading to the stimulating pads). In order to show the feasibility of the system, one scope of the project was to investigate how the connector is fused to the leads.

The most simple and reliable method of attaching the leads to the connector pads is micro-welding. One mil platinum wires, configured in line with the length axis of the connector, will be grouped together. They are all insulated from one another. This configuration is more favorable in this design application as it results in a streamlined design, which is desirable in implanted devices. The device used to fuse the wires to the pads is a Hughes Micro-Welder. This machine has the capability to perform micro-scale resistance spot welding, which is necessary to fuse the one mil wire to the connection pads and is accessible for use in the Advanced Materials Engineering Research Institute (AMERI). The benefits of welding the wires into place include the assurance of good electrical contact as well as the added pull strength of the connection. If the wires are configured in line with the length axis of the device, it is desirable for the connection of the wires to the pads to have good resistance from being torn off when pulled should the leads move around while implanted. Loss of connection between the wires and the pads would result in stimulation loss of the corresponding leads. The method of micro-spot welding, using equipment similar to the Hughes Micro-Welder, is widely seen in medical implanted devices.

# 1.3.4 O-rings. The microconnector required an O-ring design in order to prevent the entrance of an unwanted fluid to the device since this will be placed in the human body. Because it requires very little room and is lightweight, an O-ring is the preferred method of sealing the device. Possible failure of the O-ring is never sudden and could usually be recognized. Furthermore it is less likely to cause structural damage. For these reasons, an O-ring was used for our device to prevent leakage. An O-ring is a toric joint used for sealing. It closes off a pathway in order to prevent unwanted air or fluid from escaping as well as unwanted fluid from entering the device. O-rings are placed in flanges or grooves so they can block the passage. This is accomplished through material memory. When an O-ring is compressed in a space, it wants to return to its original shape. Because of this, it pushes against the walls of the groove in which it was placed, creating a seal through force.

Figure 3: O-Ring Pressure. (McGrosky, 2013)

However, it is important to keep in mind what causes an effective seal. This is determined through the inner diameter of the O-ring. The stretch of the gland diameter should not be greater than 5%, rather it should remain between 1% and 5%.

O-Ring ID: Groove Diameter% of Stretch Desired

If the O-ring has a higher stretch percentage, the pressure will cause wear in addition to cross section reduction.

For this device, the use of a micro O-ring is necessary due to the small scale of the connector. Different manufacturers were considered such as Parker O-Ring as well as Apple’s MicrOring Seals. Both manufacturers offer O-rings of very minute sizes (with Apple’s MicrOring Seals having a wider variety of choices). In order to select the optimum O-ring, the dimensions, biocompatibility, and material physical properties had to be considered. Because the O-ring is so small, it has to have a higher tolerance since there is less room for failures.

Furthermore, the O-ring must be composed of medical-grade material in order to be biocompatible with the human body and withstand the corrosive environment. The O-ring can be inserted during lamination. If the elastomer material of the O-ring is not compatible with its surroundings, the O-ring may fail, which will in turn cause the entire device to fail. In order to determine the success of the O-ring, the device will be subject to various tests such as the fine leak and gross leak tests, which are later discussed. However, due to limited resources and time, these tests were not accomplished, but in a future design, they will be.

# 1.3.5 Devices used to manufacture. The use of AMERI labs makes manufacturing devices on such a small scale possible by providing the necessary equipment and machinery. In order to use the green ceramic tape, a precision CNC punch is needed to punch the vias. The KEKO PAM Precision punch machines are able to punch very precise holes on a very small scale, between 100, 150, and 250 μm in diameter. The design is exported in a CAD file, which is uploaded into the punch machine. The machine then takes the LTCC or HTCC tape and punches out the design, which consists of the vias and the alignment holes that are used to align the layers of tape. About 25 layers of the design can be made out of one sheet of LTCC or HTCC.

Once the designs are punched out, the sheet is placed on the PTC-1000 vacuum assisted bladder filler, which filled the holes with the high-viscosity silver or platinum ink by applying downward pressure evenly across the entire sheet. After the bladder filler operation is completed, the sheets undergo a process similar to screen printing, where the ink is sheared over the vias. These two processes done in conjunction allow for a high chance of every via being filled. During the filling process, a mask is placed over the tape that allows only the vias to be exposed to the ink. The mask is made at the same time the tape is punched in the machine. This ensures that the same mechanical pins that align the tape on the machine can be used to align the mask as well.

After the vias are filled, the sheet of tape is placed in a small oven set to about 85 degrees Celsius for 10-15 minutes, solidifying the ink enough so it can stay in the vias while the tape is cut on the laser cutter. After the layers are cut out on the laser cutter, they are stacked using a simple jig consisting of four alignment posts and two metal plates. The stack is the placed in the PTC lamination press, which uses heat and about 1,500 psi of pressure to laminate the individual layers into one, which can then be fired in Lindberg 1700 furnace. The box-style oven is favorable in this application as it gives greater control over the desired temperature and uniformity of heat distribution.

Figure 4: Prototype Trial I. The blue area is the LTCC material while the white area is the HTCC. In this step, the removal of the HTCC was attempted to leave the silver pins in place.

Figure 5: Close-Up View of Microconnector. On the left-hand side of the figure, scraping of the HTCC layer was done to remove the layer while leaving the pins.

# 2. Design Alternatives

# 2.1 Overview of Conceptual Design

Before conceiving any other designs, it is important to test the manufacturability of certain components. The simplest design has been chosen due to the size of the connector and the constraints of manufacturing it. Once the preliminary testing of what can be manufactured is complete, then alternative designs can be considered. When designing alternatives, components such as size, shape, amount of pins, and additional aspects can be changed and improved; however, it is all reliant upon manufacturability of the device.

This particular connector is the testing grounds for what is possible to manufacture on such a small scale. There have not been any similar devices developed that can be used as reference to what can be accomplished. It is still uncertain whether this particular design concept of pins and holes can be manufactured, which is why preliminary testing must be conducted. The first component of the microconnector that was tested for manufacturability is the pins. If the pins can be manufactured, then this design can be further developed. In addition to making the pins, the routing of the wires that connect the pins to the pads also has to be tested. This is done by screen printing but because of the small size, it became a tough process in manufacturing. Because screen printing at such a small scale was impossible, a different approach was taken to the design. Instead of the pins connecting to the pads using the screen printing process, the pins were connected directly to the back. Once these two factors were tested for manufacturability, the final design was developed.

When designing the microconnector, several factors have to be taken into account. The major constricting factor in designing the microconnector is the manufacturing process. In addition, the ability to make the microconnector hermetic also plays a major role in the design of the microconnector. Because of these two factors, there were many limiting constraints on the design. To keep manufacturability as simple as possible, the connector consists of 32 posts and 32 holes. One half of the connector will have posts while the other half will have holes. The device connects through the posts going inside the holes. This method seems to be the simplest to manufacture. Because of manufacturing at such a small size, the major change in design occurs with the orientation of the pins and the shape of the connector.

# 2.2 Design Alternate 1

The first design considered was a connector with a 4 by 8 array of pins and holes, and pads on the other side for micro-welding the wires of the leads. This particular size of the connector is 3 mm wide and 14 mm long. This proved to be an effective size and a good pin and hole configuration. In addition, the manufacturability of this specific design was simple as well, but there were some minor problems with this arrangement. One of the concerns with this design is that it is too wide.  When dealing with implantable medical devices, the microconnector should be streamlined and as thin as possible. This is because it is easier to implant, and it takes up less of a footprint inside the human body. Shown in Figure 1 is the first design consideration. Notice the width compared to the length.

Figure 6: Preliminary Design of Microconnector.

# 2.3 Design Alternate 2

A secondary design alternative was a circular connector. This type of geometry allows for more pins to be placed, which in turn provides an improved version but has some design constraints. Furthermore, this design incorporates an O-ring, which allows for the device to be hermetic. Similar to the first proposed design, the circular geometry is not streamlined, and it is 7mm in diameter. This occupies a lot of space in terms of an implantable connector device. In addition, the pads that connect the wires from the neurostimulator to the leads will have to be placed on top of the connector which makes the connector sit in a vertical position rather than a horizontal position. Shown in Figure 2 is the circular design alternative.

Figure 7: Circular Microconnector Design Alternative.

# 2.4 Proposed Design

With considerations from the design alternatives, a proposed design was produced. Using the same considerations for manufacturability and streamlining the design, a new arrangement of the pins and holes were set up. The new arrangement features 3 rows of ten pins with two additional pins on each end in the middle. The proposed design also incorporates an O-ring for having a hermetic boundary set up. Similar to the first design alternative, pads were set up on each end of the microconnector for micro-welding the lead wires. The final width of the design is 2 mm and a length of 12 mm. Although this may not seem like a major improvement from the first design alternative, the small decrease in size has a major effect in the human body.

Figure 8: Proposed Design - Pins and O-Ring Placement.

Figure 9: Proposed Design - Closed and Sealed.

# 3. Project Management

# 3.1 Breakdown of Work into Specific Tasks

The responsibilities have been well distributed thus far.  The entire team has all gone to the lab together to get a general overview of the machines in the lab and the process involved for creating our connector in the ceramics lab. While each task will have team input, certain tasks have been divided so each individual could become the expert of a given component.  Each member put in approximately 40 hours for the fall semester including 8 hours in the lab. In addition to the lab time, we had to take three safety courses to gain access to the lab, which included Hazard Communication, Fire Protection, and Laboratory Safety.

In the spring semester, each member of the group spent around 40 hours in the lab. We were restricted from using some vital lab equipment to our project unless Dr. Karbasi was present.  Outside of the lab, Samuel and MariaMartha spent 50 hours working on all of the design modifications which included SolidWorks for the 3D models and AutoCAD for the 2-D models. AutoCAD designs were required for the machines in the lab. Eric researched the interconnect possibilities and contacted potential vendors to check for availability and costs of certain products. He also contacted Robert from B&B molders in Mishawaka, Indiana to get a quote for the mold of the housing. James initially looked into repairing the micro-welder and then turned his focus into how to machine the housing at FIU facilities.  He then 3D printed the housing to confirm if the resolution was high enough for our tiny part. The tasks were divided evenly, and everyone contributed to the project even if it was not specifically assigned to a certain person.

The group also met with Bill Box, the founder of Mark Two Engineering in Miami Lakes, to seek advice on machining the housing.  Mark Two specializes in complex, high precision machining for biocompatible metal parts. Upon discussing the design of the microconnector, Bill provided insight and advice on how to mold the device as well as how to keep it secure. Table 1 illustrates the work distribution among team members.

Table 1: Work Distribution.

| Team Member | Design | Lead Attachment | Testing | Manufacturing | Presentation |
| --- | --- | --- | --- | --- | --- |
| James Ciurdar |  | *** |  | *** | *** |
| Eric Doan |  |  | *** | *** | *** |
| Samuel Grada | *** |  |  | *** | *** |
| MariaMartha Lopez | *** |  |  | *** | *** |

# 3.2 Organization of Work and Timeline

Table 2: Project Timeline for Fall 2013.

Table 3: Updated Project Timeline from Spring 2014.

# 4. Engineering Design and Analysis

# 4.1 Structural Design

The design for the microconnector does not call for a lot of structural considerations. This device will not be subjected to immense amount of forces. The one thing that needs to be considered is the microconnector becoming undone while the device is implanted. For this reason, a titanium clip is placed over the connector. This keeps the two halves of the connector from separating. Another consideration for keeping the two halves together would be using a rubber band to tightly bind them. In addition, there is a strain relief system in place for the wires coming out of the connector. Providing strain relief to the wires will limit the amount of force being applied to the micro-welds as well as to the junction point between the wire and the connector itself. This will allow for the wires to have greater longevity since it will have less risk to being damaged easily.

# 5. Prototype Construction

# 5.1 Description of Prototype

One important aspect of the design is the connection pad design. In order for the 32, one mil, platinum wires to be easily welded into place while maintaining proper insulation and in-line configuration, there must be some separation of the pads. Separation in the z-direction would allow for the wires to be attached in groups (the group size being undetermined due to limits in manufacturability). To allow for some separation of the pads, one possibility is to manufacture the pads in steps. One way to do this would be to mill the steps after the connector halves are co-fired. The design of the steps would necessitate multiple layers of platinum ink printing, which is electrically conducting ink printed on top of the LTCC material. This ink allows for the posts and holes to be connected to the corresponding pads within the connector body. Using this step design also allows the entire connector assembly to be as thin as possible while remaining resistant to outside forces.

Below, the initial prototype design can be seen in a size comparison. This was made to determine the manufacturability of the pins.

Figure 10: Prototype Design in Comparison to a Dime.

Figure 11: Prototype of the Male Side with Silver Pins.

While the creation of the pins in this manner is unprecedented, improvement is necessary.  The pins were not very rigid. The first four pins in the second row partially snapped with minimal pressure.  The pins are 6 mils in diameter with a length of 12 mils, which is equal to 3 sheets of LTCC.  The first solution is to reduce the length of the pin to 8 mils.  The second solution is to change the material of the pin to a biocompatible platinum-rhodium, which has more than double the modulus of rigidity of silver (Koch, 2014).

# 5.2 Prototype Cost Analysis

The focus of the prototype is to determine manufacturability while trying to minimize costs.  In order to minimize costs, LTCC was used instead of HTCC. LTCC costs $20 per sheet compared to $50 for HTCC.  Each sheet could produce 25 layers with each device requiring five layers.  The pastes used to make the posts for the prototype were silver and graphite. Silver paste costs $1.25 per gram while graphite costs $7.00 per gram.  This compares very favorably to platinum paste which costs around $80 per gram. Platinum is the proposed material for the final device.  There was approximately one gram of each paste for the prototype.

The second component of cost analysis for the prototype involves the human hours spent in the lab.  As mentioned earlier, each member spent 8 hours in the lab in helping create the prototype.  In addition to group members, Dr. Ali Karbasi assisted and instructed all group members for the 8 hours completed in the lab. The breakdown can be seen in Table 4.

Table 4: Cost Analysis

| Item | Cost | Units Used | Hours | Hourly Rate | Cost Per Unit |
| --- | --- | --- | --- | --- | --- |
| LTCC | $20.00 | 0.2 |  |  | $4.00 |
| Graphite Paste | $7.00 | 1 |  |  | $7.00 |
| Silver Paste | $1.25 | 1 |  |  | $1.25 |
| Platinum Paste | $80.00 | 0 |  |  | $0.00 |
| Dr. Ali Karbasi |  |  | 8 | 80 | $640.00 |
| Senior Group Lab Time |  |  | 32 | 40 | $1,280.00 |
|  |  |  |  | Total Cost: | $1,932.25 |

Figure 11: Cost for Half of a Prototype.

# 5.3 Prototype Iterations

After failed attempts at a successful construction of both parts of the prototype, more experimental iterations of the manufacturing process were needed. There was a recurring problem in the making of the male portion of the connector. The HTCC Alumina was used in the prototype as the sacrificial material to form the pins and was not behaving as expected. In theory, the HTCC, when fired at sub-sintering temperatures, was expected to brush away after firing, thus leaving the formed silver pins left behind, which would then be inserted into the female portion of the connector. However, the HTCC reacted with the silver and did not brush away the first couple of times it was attempted.

The construction of these pins is one of the most vital portions of the connector, and the successful construction of the pins was the most experimental part of the device. Not knowing whether or not the pins could be successfully formed placed the entire design of the connector in jeopardy as it is the basis for the mechanical and electrical connection. It is believed that the first HTCC material that was used failed to disintegrate as expected because of the close proximity between the pins, which may have allowed the glass particles present in the ceramic HTCC material to sinter at a lower temperature.

Being aware of this, another 96% Alumina HTCC material was chosen because it has less of the glass particles which help it sinter in hopes that the HTCC would disintegrate as expected when co-fired. However, all attempts at using HTCC as the sacrificial material failed. The solution to this problem was to use a form of carbon tape as the sacrificial material instead of the HTCC as it is much more likely to be burned off when fired in the furnace. This method of construction was successful in creating the pins, leaving the surface of the hardened LTCC material clean and ready for testing.

Figure 12: Prototype Design with Carbon Tape as Sacrificial Material.  Pins were successfully manufactured.

Figure 13: Close Up View of Silver Pin in LTCC Layers.

# 6. Design Considerations

# 6.1 Design Adaptation due to Manufacturability

When manufacturing and testing the proposed design, several problems arose, which changed the direction of the design. Manufacturing the male and female sides of the microconnector was successfully completed in the AMERI lab. However, creating additional features to the layers of the ceramics such as screen-printing conductive ink on the different layers adds more difficulty to the manufacturing process. In order to keep the manufacturing process simple for the ceramic male and female halves, the pads were transferred to the back of the ceramic parts. This allowed for simple punching and filling to be done and avoid the screen-printing process. The images below show how the welding pad section on the device was simplified.

Figure 14: Former Design using Screen Printing. The pins are going to pads using the screen printing process.

Figure 15: Male Half of New Design. This drawing shows the pins on the male half going directly to pads on the back of the ceramic part.

Figure 16: Female Half of New Design. The vias connect directly to the back of the ceramic parts as new design.

In addition, because the device is extremely small, trying to add an O-ring to the ceramic section was very complex. Too many factors needed to be accounted for such as shrinkage, alignment, and the layering process. This problem caused us to create a housing for the ceramic halves to be placed in. The housing would be made out of a biocompatible material that would be able to withstand the harmful conditions inside the human body. In addition, the housing would now hold the O-ring. Although there would be an increase in size, it is still smaller than the considered designs for the V.A. Clinic.

Figure 17: Former Design with the O-ring. The location of the O-ring can be seen in the ceramic part of the microconnector.

Figure 18: Housing Created and Ceramic Part Placed in the Housing. The O-ring is located in the housing and is not in the ceramic part anymore.

The final problem that not having a housing caused was the alignment of the pins. The pins of the device are less than 10 mils in diameter, which makes them ductile and difficult to align at such a small scale. Adding the housing allowed for alignment pins to be created. The alignment pins would be molded on the device. One half would have the pins and the other would have the slots that would allow perfect alignment of the device.  The images shown below illustrate the alignment pins.

Figure 19: Isometric View of Housing Shown with Alignment Pins

Figure 20: Top View of Alignment Pins in Housing.

Figure 21: View of the Housing with Alignment Pins Before Connection.

Moreover, many options were looked over in order to manufacture the housing. While exploring more cost-effective options, rapid prototyping was considered. Due to limited funding, the housing could not be made with the appropriate materials. FineLine Ptrototyping provided a quote for building the housing with a material similar to ABS. DSM Somos Bioclear was the intended material to build a housing using a high resolution stereolithography. The total price for only one half of the housing was $201, which was far above what our funds could cover. Furthermore, most companies accessible to us, such as Stratasys or HK3D Printing, can only print biocompatible materials that are subjected to be either in the mouth or on the skin for a prolonged amount of time. None of them could print using a material that would actually be implanted inside the human body. However, to rapid prototype using biocompatible, non-toxic material is possible. According to KurzweilAI, the discovery of using riboflavin, also known as vitamin B2, made it feasible to rapid prototype a material that is safe for implantation. Researchers at the University of North Carolina use the two-photon polymerization (2PP) technique for 3D printing a variety of small, biocompatible devices that can be applied to tissue engineering, implantable devices, wound dressings, and stem cells. Therefore, it is possible to 3D print the housing of the microconnector in a biocompatible material.

Because this technique is not easily accessible to our team since research on this technique is still being completed, we decided to 3D print the housing in non-biocompatible material to be able to test the ability of 3D printing it. We were able to determine that the housing can indeed be made in a biocompatible material because ultimately, the part would be built using a mold. By using a mold, the biocompatible material can be made out of polyether ether ketone (PEEK), which is a material highly used in medical spinal implants. Due to this fact, we printed the housing using the MakerBot Replicator 2Desktop 3D Printer. One half of the housing without the alignment pins was printed using the MakerBot. The layer resolution of the printer is 100 microns. The material used was a biodegradable polymer called polylactic acid (PLA). This PLA filament is resistive to shrinking, low in cost, and provides a durable product. Two housings were printed; one being the actual size of the housing and the other was scaled up. The housing that was scaled up is clearly more well-defined. Due to the dispenser head’s low resolution, the actual size of the housing does not have a resolution as high as the bigger model. While this poses a problem because of the 3D printer that was used, it is possible to 3D print micro-materials because there are micro-dispensers that are currently available in the market. The following figures show the 3D printed housing halves.

Figure 22: Larger 3D Printed Half of Housing.

Figure 23: Size Comparison of 3D Printed Housings. The housing that shows the actual size is not as defined as the bigger half.

Figure 24: Measurement of Housing. The housing that was 3D printed to scale is shown to be about 1.5 mm long including the strain relief.

Many of the problems that we had with the manufacturability of certain features on the ceramic part were fixed by creating a housing. The housing also provides a better surface in which the doctors can perform the surgery easier. Shown in the following figures are images of the fully assembled microconnector.

Figure 25: Isometric View of the Microconnector. This is how the microconnector would look fully assembled.

Figure 26: Top View of Microconnector.

Figure 27: Front View of Microconnector.

Once the housing was prototyped, there were several issues that had to be addressed. One of the major issues was the strain relief not adding enough flexibility to provide the adequate relief of the wires. The second problem that was found was the grooves in the housing to provide an indent for the two clips holding the two halves together. These changes to the housing did not affect changes in dimensions. In addition to the separate strain relief and grooves, the alignment posts were changed to slots to allow for simpler manufacturability and installation of the device. Shown in the following images is the final design of the microconnector.

Figure 28: Addition of Strain Relief to the Housing.

Figure 29:  Fully Assembled Female Half of the Microconnector.

Figure 30: Fully Assembled Male Side of Microconnector.

Figure 31: Fully Assembled Front View of Device.

Figure 32: Fully Assembled Isometric View of Part.

Figure 33: Exploded View of How Microconnector will be Assembled.

Figure 34: Major Dimensions of Microconnector.

# 6.2 Interconnection Options

The very center of the project has the mechanical connection between the male pins and the female vias, which provide the electrical connection. While the majority of the project definition resides in the proof of concept, which is the formation of the pins in thick film technology, a critical part of the desired end result for the project is the ability to use these pins to provide the connection. While it may be possible to try to form the electrical connection by bottoming out the pins in the vias, therefore making the connection to the conductive material present in the vias, it is favorable to include another method of connection as the unassisted method does not leave much room for error. This error includes possible irregularities in the length of the pins, depths of the holes, and unequal applied clamping pressure. To offset the chance of not making a good electrical connection in every via, various methods of connection were considered.

A first method for an interconnect technology is dispensing a conductive elastomer into each via. This elastomer would be a base material of some type of silicone, giving flexibility to its application. This silicone material would have a conductive material infused into its structure, preferably Platinum. The conductive particles would be dispersed within the material such that when the material has a uniaxial load applied to it, the conductive materials align and conduct the electrical signal between the pins and vias.

The second option would involve a product known as a Fuzz Button. Fuzz Buttons are flexible interconnects manufactured by Custom Interconnects.  Fuzz Buttons have been around for more than 50 years.  They miniaturized them more than 20 years ago to meet the rapidly growing microelectronic market. The Fuzz Buttons that were purchased were made of Gold-plated Beryllium Copper wire since they were the cheapest; however, the implantable device would need to use Gold plated Tungsten or Molybdenum wire since these metals are known to be biocompatible.  The wires are bundled together in a flexible spring-like structure. This is favorable as an interconnection method due to its flexibility as a mechanical structure and its reliability as an electrical connection. Fuzz Buttons come in various sizes, ranging from diameters of .010” to .500”, and lengths ranging from .020” to .500”.

Figure 35: Fuzz Button Under Magnification.

Figure 36: Fuzz Button Compared to Dime. The Fuzz Button is shown on a dime near the ear. The Fuzz Button was circled for a clearer reference.

Between the two options for an interconnection method, the Fuzz Buttons are more favorable in theory because they have a strong, yet flexible mechanical structure. The spring-like structure gives a more reliable connection between the two contact points, which is a desired trait of the project description.  They also have a history of reliability as they are used in several military and defense applications such as missiles, satellites, and in high end RF products.  In order for the Fuzz Buttons to work, they first have to be modified because the length of the 10 mil diameter Fuzz Button is 40 mils.  For our application, we need the length to be less than 15 mils.  The diameter of the Fuzz Button also required us to redesign the microconnector. The conductive elastomer did not require us to redesign our device; however, we learned there currently is no commercially available conductive elastomer available for a long term implant though there is one in development from NuSil, according to David Mitchel a regional salesperson for NuSil.

Figure 37: Fuzz Button Comparison to Female Vias.  This shows the 10 by 40 mil Fuzz Button and 20 by 20 Fuzz Button on top of the female side of our connector.

The microconnector was redesigned to make the Fuzz Button a viable option.  The size of the via diameter was changed to 400 microns, and the pitch was altered to 800 microns.  In order to make the holes larger, a laser cutting machine was necessary, as the punch machine did not have a punch head larger than 250 microns. To obtain the desired clearance fit, the holes were made to 400 microns as a trial with an 800 micron pitch. This sizing would theoretically allow the Fuzz Buttons to fit and leave enough material between the holes to have a stable structure. However, it was discovered that when making holes in the tape using the laser, the Mylar backing of the tape melts and adheres to the HTCC tape, leaving behind a torn sheet of HTCC when peeling off the Mylar after filling the holes with the conductive material before firing.

Furthermore, the filling process did not respond well to the new sizing, as the material did not stay in the holes after undergoing the bladder filler operation and shearing operation with the squeegee. We then attempted another modified design using a 320 micron diameter hole using the laser.  This would be the minimum size we could manufacture to utilize the Fuzz Button.  The HTCC did not rip this time during the fill process; however, the same problems occurred with the Mylar. Therefore, the new sizing practically does not work. In order for this sizing option to work, the laser problems would have to be remedied by lowering the power, and a new filling process would have to be used.

The last potential solution to this problem attempted was to use the original size hole of 250 microns for the male part and 380 microns for the female part, and then add another step at the end of the process.  After the female side is laminated and fired, we drilled the holes using the router. This solution ended up being successful; the Fuzz Button was able to fit in the female part of the ceramic. It was able to stay in the hole by using a silver paste to adhere it inside. This can be seen in Figure 38.

Figure 38: Male and Female Half with Fuzz Button. The inserted Fuzz Button can be seen circled in the female half of the ceramic part.

# 6.3 Environmental Impact

For our part, we used ceramic material to make the female and male halves. Ceramic is not biodegradable. However, our waste from the project is minimal as seen from the picture below. While this waste is minimal, in order to reduce it, we could manufacture three or four microconnectors from the size that is seen in the figure. This would reduce the amount of material being thrown away. Because we are just constructing the design as a proof of concept, minimizing waste material was not a priority since it is very minimal due to the size of the connector.

Figure 39: Waste Material Compared to a Quarter.

# 7. Testing and Evaluation

# 7.1 Overview

In order for our device to be utilized, it will need to pass several tests to determine whether it is safe and effective for use in the human body. The Food and Drug Administration (FDA) prefers testing protocols set forth by the military. The MIL-STD 883 is the military testing protocols for microelectronics.  There are over 50 tests composed of Electrical, Mechanical, and Environmental aspects of the device.  For our device, we consider five of the tests.  We also would test the durability of the pins, which is not covered from the MIL-STD test methods. In a future design, the male and female sides of the connector would have to be attached and unattached numerous times to ensure the pins did not break off.

The electrical test for our device is the test method 3014, which is the functional test.  The six specific components of this test include test temperature, power supply voltage, input voltage levels, input and output logic patterns, output threshold, and voltage minimum and maximum.  Our focus is the input and output voltage levels since our device does not generate electricity; it merely transfers signals from the stimulator to the leads.  We need to confirm we have a good connection with the pins so we are not losing any voltage in our device.

The most critical test will be test method 1014.  Test method 1014 is the seal test and is composed of two components, the fine and gross leak test.  This test is primarily checking the seal of the O-ring to ensure our device is hermetic and that no moisture is entering the device potentially causing a short. The fine leak test must be done first because the liquid used in the gross leak test could fill microscopic cracks in the device potentially giving us inaccurate results.  This test involves placing the device in a chamber and then bombing it with Helium. Afterwards, the amount of Helium found inside the device is measured. This determines whether the device is hermetic since Helium is not naturally found in the atmosphere.  The maximum leak rate for a biomedical device is 1 ×10-9 atm cc/s He.  The leak rate is determined from the following equation:

From the equation above, the variables t1, t2, and Pε are determined based on the volume of the package. This is illustrated in Table 5.

Table 5: Fixed Conditions for Fine Leak Test

The “L” is the equivalent standard leak rate in atm cc/ s air while the part of the equation MaM  is a constant since this is the molecular weight of air divided by the molecular weight of helium.  The open space of our device is well below .05 cubic centimeters so the test conditions from the first row would be used.

The gross leak test can be done simply by submersing the device in a beaker of water for a couple hours. The device is then removed from the water and opened to check for any water that may have leaked into the device.  Once this test was passed, then military standards were employed by placing the device in a pressure chamber with a perfluorocarbon fluid with certain properties.

Several other tests would need to be completed in order for this device to be considered by the FDA.  A simulated body fluids test would be done to confirm that there is no corrosion on our device from the harsh environment within the human body.  This test is less of a concern as we are using materials with a long history of being biocompatible.

# 7.2 Final Design Cost Evaluation

In order to evaluate what would be the total cost of a part that would be manufactured, the projected materials that would be utilized in addition to the other things needed to manufacture were estimated. The total cost of the housing as well as the ceramic part is illustrated in the pie chart below.

Figure 40: Projected Cost of Microconnector.

# 8. Conclusion

# 8.1 Conclusion and Discussion

With increasing advancement of micro-technology as well as new discoveries in diseases, the demand for improved neurostimulators has increased. According to the American Association of Neurological Surgeons, up to 50,000 neurostimulators are implanted into patients worldwide each year. A microconnector that improves the interchangeability of the device is an effective measure against the problems with hardwired neurostimulators today. With double the amount of channels and reduced size, a streamlined design for the microconnector can accomplish this. The initial microconnector was built out of LTCC in order to reduce cost; however, for the final design, the microconnector will be composed of HTCC, titanium, and other biocompatible materials.

Using the equipment in AMERI, the microconnector can be manufactured, but due to the restrictions the size of the microconnector has, many limitations on the design were discovered. Even though certain machines and materials were unable to be applied in this design, there is technology available to make a biocompatible connector of a micro size. Although building a completely biocompatible device was not accomplished, many discoveries were made about fabricating such a device which will help pave the way for future research. Knowing the limitations of the device is a huge step in future fabrications of a microconnector. Overall, improving the connector by increasing the posts and holes will make a surgeon’s job easier as well as enhance comfort for the patient.

# 9. References

Dimensional Considerations. (2013). Retrieved February 13, 2014, from http://www.applerubber.com/seal-design-guide/oring-basics/dimensional-considerations.cfm

Example Applications of Microwelding Industry. (2013). Retrieved March 4, 2014, from http://www.tjsnow.com/microwelders/index.htm

Karbasi, A. (2012). Developing a high density pt/alumina hermetic feedthrough (Unpublished doctorate thesis). Florida International University. Retrieved February 19, 2014, from http://digitalcommons.fiu.edu/cgi/viewcontent.cgi?article=1775&context=etd

Keko equipment. (2013). Retrieved January 28, 2014, from http://www.keko-equipment.com./

Koch, M. (2014). Elastic properties of Pt, Rh, Ir and their alloys. Retrieved March 30, 2014, from http://heraeus-pt-labware.com/en/materials/dphwch/elastische_eigenschaften_/ elastische_eigenschaften__1.aspx

MakerBot Replicator 2 Desktop 3D Printer. (2013). Retrieved March 20, 2014, from http://store.makerbot.com/replicator2

McGrosky, D. T. (2013, February). Introduction to O-rings. Retrieved January 28, 2014, from http://www.satoriseal.com/technical/technical_articles/introduction_to_o-rings.htm

Medtron specify implant manual. (2008). Retrieved January 14, 2014, from http://professional.medtronic.com/wcm/groups/mdtcom_sg/@mdt/@neuro/documents/documents/scs-ld565-impmanl.pdf

Medtronic. (2010). Surgical lead impantation guide [Manual]. Author. Retrieved February 18, 2014, from http://professional.medtronic.com/wcm/groups/mdtcom_sg/@mdt/ @neuro/documents/documents/scs_specify_surg_impl_gd.pdf

Medtronic Neuromodulation. (2007). Technical design summary [Manual]. Author. Retrieved March 5, 2014, from http://professional.medtronic.com/wcm/groups/mdtcom_sg/ @mdt/@neuro/documents/documents/scs-ld565-tech-sum.pdf.

Micro Resistance Welding. (2011). Retrieved February 11, 2014, from http://www.a-i-t.com/information/micro-resistance-welding

Mihov, D., and Katerska, B. (2010). Some biocompatible materials used in medical practice. Trakia Journal of Sciences 8, 119-125.

Section I Introduction [Introduction]. (2001). In Parker O-ring handbook (50th ed., pp. 1-2-1-7). Lexington, KY: Parker Seal Group, O-Ring Division.

Spinal Cord Stimulation. (2008, October). American Association of Neurological Surgeons. Retrieved April 3, 2014, from https://www.aans.org/Patient Information/Conditions and Treatments/Spinal Cord Stimulation.aspx

Vasudev, A., Kaushik, A., Jones, K., and Bhansali, S. 2012. Prospects of low temperature co-fired ceramic (LTCC) based microfluidic systems for point-of-care biosensing and environmental sensing. Springer.

3D-printing non-toxic biocompatible medical implants. (2013, October 25). KurzweilAI. Retrieved March 30, 2014, from http://www.kurzweilai.net/3d-printing-non-toxic-biocompatible-medical-implants